## Slide 1
Quantitative Thermoacoustic Imaging using a
Clinical Ultrasound Transducer Array
S.K. Patcha, D. Hullb, W.A. Seec and G.W. Hansond
aDept. of Physics, UW-Milwaukee
bAvero Diagnostics
cDept. of Urology, Medical College of Wisconsin
dDept. of Electrical Engineering, UW-Milwaukee

## Slide 2
+ homogeneous initial conditions
Thermoacoustic Imaging
inverse source problem
- rapid heating (EM ala MRI)
- pressure jump (US)
- spherical sinogram (CT)
ns t2
ns t1
transducer
prostate
MRSI based on citrate suppression (Narayan & Kurhanewicz ‘92)
in vitro NMR more specific than PSA (Serkova 2008).
mechanical,
thermal)
thermal & mechanical
contrast = f(dielectric,
dielectric
Apply E so that
s ~ ionic content
|E| & I(t) are known, |E| homogeneous over 6 cm FOV
Ionic content of prostatic fluid negatively correlated to disease state (Kavanagh 1985).

## Slide 3
O-scope (4-chan)
stepper motor
VSX box + host
Δθ = 1.8, Δz = 0.3 & 20 mm, 90 min
acoustic couplant: 15 g/L glycine in DI water
2.25 MHz focused & heavily damped single
elements + 54 dB preamp, 128 signal aves
TA pressures weak, 1-10 Pa
P4-1 cardiac array w/96 channels, amplify
43.5 dB, 1024 aves
known freq response
TE10 waveguide

## Slide 4
Multi-Channel Viz
Acquired by 96 channel transducer array + Verasonics
slice 51 of 162 total
verumontanum
apex
base
urethra
Visualize structures not typically seen on US, CT, or MRI.

## Slide 5
Prostate #2 of 12
viz in 3D Slicer
axial
sagittal
coronal
R
L
P
A
A
P
veru
engorged vesicles
streak artifact

## Slide 6: Single element vs. clinical array
single element P4-1
combined

## Slide 7
qTA of saline, s ~ 1 S/m
1. reconstruct dp in Pa from V/Pa response
2. infer s using G = 0.1 and |E| = 6 kV/m
Pa
Pa
Pa
P4-1 V306 ave.

## Slide 8: Quantitative TCT take-aways
Acknowledgements
Clinical support: MCW CTSI & Tissue Bank, H Engelke & Dr. K Jacobsohn
Verasonics hardware: UWM Instrumentation Grant
Data acquisition: YM Qadadha, I Elsharef & M Yang
Benchtop TCT system often detects anatomical structures.
Very low US frequencies (20 kHz) critical for “qTA” imaging.
Very high (100 MHz) EM-induced thermoacoustics can provide
whole organ images over 6+ cm FOV
image of electrical conductivity, ionic content + relaxation

## Slide 9: EM irradiation power spectrum vs. k-space
S
p =
- incident pulse, V(t) = Vo I1/2(t) cos(wot)
- I1/2
power spectrum
Vo2 FI
limits resolution
limits contrast

## Slide 10
Prostate #3 of 12
axial
sagittal
coronal
R
L
P
A
S
I
R
L
A
P
ductal structure clearly visible

## Slide 11: Resolution Improvement – in-plane
80 micron copper wire phantom suspended near isocenter
2.5 cm horiz. (current induced)
10 cm vertical (TA signal)
Ideal image is an 80-micron “dot.” Actual image is called the point spread function (PSF).
Expect anisotropy from polarization

## Slide 12
SE P4-1
“old” SE “new” P4-1
X-bar || E
X-bar || E
X-bar || E
X-bar || E
X-bar E
X-bar E

## Slide 13: Best and worst-case resolution
parallel perpendicular

## Slide 14: Backup Slides

## Slide 15
Upgrade to Clinical Array + Verasonics Box
single-element transducer
lid
P4-1 clinical
array, recessed
insertion direction
rotation/ translation axis
copper EM output probe
6 cm
1 5/8” line section
64 cm
copper EM input probe
50 dB slugs
E-field polarized horizontally
Propagate ~25kW TE10 pulse
carrier freq 108 MHz
pulsewidth 700 ns

## Slide 16: k-space and image quality
log10 scale
background shading
SE P4-1 array
piezos high-pass
- P4-1 improves resolution
- contrast/accuracy suffer
miss k-space center
pixel values wrong

## Slide 17: Low Frequency Results 2.25 MHz single element transducer
display window [-0.7, +0.7] mV
verumontanum
#1
#2
#3

## Slide 18: TCT vs Gleason (primary) N=11 ROIs shown above

## Slide 19: TCT vs Gleason (primary) N=45 ROIs in posterior peripheral zone

## Slide 20: Limitations – temperature and incident power variations
ROIs posterior of each slice from 25 control specimens (hollow)
#1 #2 #3 3 annotated specimens (solid)
average TCT vs. temp

## Slide 21: Limitations – incident power
P1inc  incident
P1ref  reflected
P2inc ☐ transmitted
line section & 50 dB atten. slugs by Bird Electronics
hollow 2-channel data
w/o reflected power
unloaded (w/o prostate)
loaded
loaded
loaded
solid 4-channel data
monitor reflected power

## Slide 22: Preliminary Conclusions
- observed suppressed TA signal from diseased PZ tissue.
- VHF heating yields TA contrast driven by ionic content.
(for 45 ROIs from only 3 specimens 0.4 mV provides
100% sensitivity but 29% specificity.)
Controls for Limitations
5°C temp range likely contributed to signal difference,
BUT not enough to account for observations
EM incident power greatest for prostate #1, which produced least signal
Acknowledgements MCW Support Team M Thomas - CTSI funded (NIH #8UL1TR000055) H Kelly, H Engelke
students: E Hanson (McGill)

## Slide 23
+ homogeneous initial conditions
Thermoacoustic Imaging
inverse source problem
- rapid EM heating (MRI)
- pressure jump (US)
- spherical sinogram (CT)
ns t2
ns t1
Ionic content of prostatic fluid negatively correlated to disease state (Kavanagh 1985).
thermal)
thermal
mechanical,
mechanical
contrast = f(dielectric,
electrical
transducer
prostate
MRSI based on citrate suppression (Narayan & Kurhanewicz ‘92)
in vitro NMR more specific than PSA (Serkova 2008).